## Tabelle1
Flows
Point of measurement | Flow TSI1 | Flow TSI2 | Flow Bios 220 | Flow Bios 220 big
Total inlet |  |  |  | TSI big: 17.6
APS | 3.66 | 3.84 |  | 3.68
WIBS | 0.33 | 0.31 | 0.3
SMPS | 0.33 | 0.33 | 0.22 | 0.31
SMPS Makeupflow | 0.58
CPC total | 0.31 | 0.32 | 0.31
CCNC | 0.46 | 0.49 | 0.4 | 0.49
AMS | 0.07 | 0.08 | 0.07
Aethalometer | 2.05 | 2.14 |  | 2.03
Total pump |  |  |  | 10
SUM | 17.79
Interstitial Inlet |  |  |  | TSI big: 16.64
SO2 | 0.45
Picarro | 0.23
O3
AMS | 0.7
CPC INT | 0.29
CCNC | 0.47 |  | 0.46
SMPS | 0.29
Aethalometer | 2
Makeupflow SMPS | 0.58 |  |  | 0.59
Total Pump Int | 11.23 |  |  | TSI big: 11.23
Sum: | 16.24

## Tabelle2
Logbook inlet Particle counter
Folder for CPC 2025 measurements: Data/CPCDMAlogger/inter/InletCharacterization |  | 2019-08-16
1 |  | Set valves to total inlet
2 |  | CPC3776 on total inlet
3 |  | Select 20nm particles with DMA
4 |  | Connect Nebulizer with MiliQ
5 |  | Measure Flow: 13.6 lpm on big filters, 0.62 lpm on aerosol inlet
6 |  | Compare both CPC at inlet: both measure ca. 180 p/cc
7 |  | Measure Zero: Disconnect aerosol flow and set filter in front: 2-4 p/cc
8 |  | Total Inlet Flow: ca. 9lpm
9 |  | Interst. Inlet Flow: ca. 11lpm
10 |  | CPC3776 on zero: 0-10 p/cc. On container air: 250p/cc. On container air with SMPS in front: 15 p/cc
11 |  | Software CPCforcalibration is for DMA size selection.
12 |  | Sample Flow with nebulizer (miliQ), DMA on 20nm: 0.975 lpm
13 |  | Main Flow: 13.05
14 | 1655 | Flush with MiliQ
| 1704 | Change PSL: 46nm
| 1705 | Change Settings on SMPS to 46nm.

## Tabelle3
Inlet Characterization
Path to files:
CPC 3010 | C/Data/CPCDMAlogger/total/InletCharacterizationDMA |  |  | Sample flow: where? | 0.8 | With this, the aerosol concentration goes down to 46 p/cc
CPC 3776 | C/Data/CPCDMAlogger/InletCharakterization_3776 |  |  | Makeup Flow where? | 15.7
CPC 3025 | C/Data/CPCDMAlogger/Inter/InletCharacterization |  |  | New sample Flow: whats that? | 1lpm | Concentration for 31nm: 1220 p/cc
|  |  |  | Delay time: Approx. 10s |  | Sheath flow DMA: | 3lpm
3776 shows approx. 50p/cc less on the vi than in the display
31nm PSL Test 19.8.2029
| Time | Concentrattion 3776 at total inlet | Concentration 3025 | Remarks | Points 3025 | Points 3776
CCNC | 15:07 | 244 | 232 | stop 15.21 | 792-853 | 73491 - 74101
Zero | 15:22 | 3 | 6
SMPS | 15:43 | 290 | 260 | 16:05 | 1009-1129 | 75658 - 76852
Zero | 16:08 | 3 |  |  | 1151-1157 | 77084-77140
AMS | 16:12 | 320 | 280 | stop 16:22 | 1194-1231 | 77504 - 77876
Zero | 16:24 | 3 | 2 | stop 16:26 |  | 78034 -78342
Aethalometer | 16:36 | 360 | 380 | Stop 16:46 | 1324-1379 | 78788 - 79349
MiliQFlush | 16:50 |  |  |  |  | 79730 - 80284
WIBS
APS
46nm PSL Test 19.8.2019
| Time | Concentrattion 3776 at total inlet | Concentration 3025 | Remarks
CCNC | 17:10 | 330 | 330 | stop: 17:22. | 1527-1595 | 80794 - 81472
Test on total inlet | 17:23 | 340 | 380
Aethalometer | 17:25 | 370 | 390 | Display values | 1623-1726 | 81735 - 82778
SMPS | 17:49 | 400 | 420 | stop: 18:01 | 1753-1835 | 83052 - 83869
Zero | 18:02 | 2 | 1.55 | stop 18:15 | 1852-1907 | 84052-84593
AMS | 18:18 | 410 | 390 | stop: 18:29 | 1936-1995 | 84887 - 85477
WIBS
APS
152nm PSL Test: 20.8.2019
| Time | Concentrattion 3776 at total inlet | Concentration 3025 | Remarks
CCNC | 9:37 | 112 | 131 | stop: 09:49 | 3302-3361 | 33600-34195
SMPS | 9:14 | 135 | 150 | stop: 09:35 | 3159-3281 | 32174-33392
Check | Check error on 3776: | it is countinfg ca. 20% too low.
AMS | 13:20 | 130 | 118 | stop: 13:33 | 3996-4054 | 37440-38025
Aethalometer | 0.35208 | 100 | 116 | stop: 08:40 | 2882-2946 | 29391-30037
Aethalometer | 13:35 | 124 | 135 | stop: 13:47 | 4081-4142 | 38327-38930
Flush with MiliQ | 13:48 | x | x | stop: 13:53
WIBS
APS
508nm PSL Test: 20.8.2019, start DMA: 13:53
| Time | Concentrattion 3776 at total inlet | Concentration 3025 | Remarks
Concentrations are very low. Turn up nebulizer flow to max (aerosol flow is now 1.98lpm). Total inlet shows around 15 p/cc. Increase integration time for both cpc's to 20s.
Zero at CCNC | 15:26 | 4 | 2.9 | stop: 15:33 | 4648-4691 | 43094-43137
CCNC | 15:34 | 190 | 106 | stop: 15:54, concentration was steadily increasing, from 190 at the beginning to 340 at the end | 4710-4800 | 43157-43247
SMPS | 16:09 | 435 | 330 | stop: 16:22, total conc.: 490, SMPS conc. 375 | 4919-4971 | 43374-43427
AMS | 15:56 | 350 | 283 | stop: 16:08, total: 411p/cc, AMS: 315p/cc | 4830-4892 | 43279-43344
Aethalometer | 15:06:00?? | 29 | 31 | stop: 15:18
Aethalometer | 16:24 | 510 | 434 | stop: 16:32 | 4990-5031 | 43452-43493
flush MiliQ | 16:35
WIBS
APS
345nm PSL Test: 20.8.2019, start DMA:
| Time | Concentrattion 3776 at total inlet | Concentration 3025 | Stop
CCNC | 17:06 | 140 | 96 | 17:18 | 5266-5307 | 43730-43770
SMPS | 17:19 | 140 | 90 | 17:30 | 5324-5375
AMS | 17:31 | 138 | 100 | 17:45
Aethalometer | 16:54 | 140 | 117 | 17:04
WIBS | 17:50 | 135 at WIBS | x | 17:55
APS

## Tabelle4
PSL (nm): | 30 |  |  | 45 |  |  | 150 |  |  | 350 |  |  | 508
CPC: | 3776 | 3025 | quotient | 3776 | 3025 | quotient | 3776 | 3025 | quotient | 3776 | 3025 | quotient | 3776 | 3025 | quotient
CCNC | 238 | 234 | 0.98319 | 292 | 353 | 1.2089 | 113 | 131 | 1.1593 | 117 | 95 | 0.81197 | 207 | 128 | 0.61836
SMPS | 256 | 280 | 1.0938 | 350 | 430 | 1.2286 | 123 | 146 | 1.187 | 116 | 88 | 0.75862 | 395 | 331 | 0.83797
AMS | 279 | 276 | 0.98925 | 357 | 397 | 1.112 | 109 | 125 | 1.1468 | 116 | 102 | 0.87931 | 318 | 286 | 0.89937
Aethalometer | 315 | 403 | 1.2794 | 337 | 416 | 1.2344 | 104 | 123 | 1.1827 | 113 | 113 | 1 | 426 | 414 | 0.97183
Inlet directcomp |  |  |  | 306 | 395 | 1.2908
|  | 'Instr' | '30nm_76' | 'stdev30_76' | '30nm_25' | 'stdev30_25' | '45nm_76' | 'stdev45_76' | '45nm_25' | 'stdev45_25' | '150nm_76' | 'stdev150_76' | '150nm_25' | 'stdev150_25' | '350nm_76' | 'stdev350_76'
|  | 'CCNC' | 238.8 | 14.8 | 234.8 | 9.1 | 292.8 | 17.4 | 353.6 | 12.4 | 113 | 10.7 | 131.7 | 5.6 | 117 | 5.1
|  | 'SMPS' | 256.2 | 16.8 | 280.9 | 12.5 | 350.4 | 18.7 | 430.8 | 10.1 | 123.1 | 10.9 | 146.2 | 7.1 | 116.4 | 4.8 | '350nm_25' | 'stdev350_25' | '500nm_76' | 'stdev500_76' | '500nm_25' | 'stdev500_25']
|  | 'AMS' | 279.4 | 15.5 | 276.1 | 9.6 | 357.7 | 17.8 | 397.7 | 9.6 | 109.3 | 10.3 | 125.1 | 5.8 | 116.3 | 7.1 | 95.8 | 5.2 | 207.5 | 37.6 | 128.6 | 24.4
|  | 'Aeth' | 315 | 18.8 | 403.3 | 12.2 | 337.7 | 17.4 | 416 | 11.1 | 104.6 | 10 | 123.3 | 4.8 | 113 | 6 | 88.9 | 5.4 | 395.5 | 27.9 | 332 | 25.8
[17 empty] | 102.4 | 7.8 | 318.7 | 17.7 | 286.9 | 18.1
[17 empty] | 113.7 | 6.5 | 426.5 | 28.7 | 414.7 | 26.8
|  |  | q30 | stdev30 | q45 | stdev45 | q150 | stdev150 | q350 | stdev350 | q500 | stdev500
|  |  | 30 |  | 45 |  | 150 |  | 350 |  | 500
|  | ccnc | 0.98325 | 0.071872 | 1.2077 | 0.08333 | 1.1655 | 0.12098 | 0.8188 | 0.057002 | 0.61976 | 0.1626
|  | smps | 1.0964 | 0.086888 | 1.2295 | 0.071665 | 1.1877 | 0.11994 | 0.76375 | 0.056072 | 0.83944 | 0.088103
|  | ams | 0.98819 | 0.064698 | 1.1118 | 0.061493 | 1.1446 | 0.12021 | 0.88048 | 0.08595 | 0.90022 | 0.075665
|  | aet | 1.2803 | 0.085667 | 1.2319 | 0.071478 | 1.1788 | 0.12168 | 1.0062 | 0.078506 | 0.97233 | 0.090717